## Worksheet 1
Error | Type | Standard | Sensitivity | Uncertainty
source |  | uncertainty | coefficient | contribution
xi |  | u(xi) | c(xi) | u(xi)⋅c(xi), µm
Calibrated length of the RWP at 20 ∘C
(l20 rwp=125.0000 mm) | B | 0.1 µm | 1 | 0.1
Measured length of the WP
(lwp=125.0048 mm) | B | 1.4 µm | 1 | 1.4
– random effects during measurements | A | 0.9 µm
– measuring instrument | B | 1.0 µm
Measured length of the RWP
(lrwp=125.0043 mm) | B | 1.2 µm | −1 | −1.2
– random effects during measurements | A | 0.6 µm
– measuring instrument | B | 1.0 µm
Measured temperature of the RWP
(trwp=Δtrwp+20 ∘C = 35 ∘C) | B | 1.3 K | 0 µm K−1 | 0
– random effects during measurements | A | 1.2 K
– measuring instrument | B | 0.5 K
Calibrated CTE of the RWP
(αrwp=10.52×10-6 K−1) | B | 0.6×10-6 K−1 | 0 µm K | 0
Known/assumed CTE of the scale
(αrwp=8×10-6 K−1) | – | Can be neglected | 0 µm K−1 | 0
Possible temperature difference between the RWP and the scale during measurements of length of the WP
(δtscwp=0 K) | B | 0.6 K | 1.0 µm K−1 | 0.6
Possible temperature difference between RWP and the scale temperature during measurements of length of the RWP
(δtscrwp=0 K) | B | 0.6 K | −1.0 µm K−1 | −0.6
Possible temperature difference between the RWP and the WP
(δtwp=0 K) | B | 0.6 K | −1.3 µm K−1 | −0.8
Possible CTE difference between RWP and WP
(δαwp=0×10-6 K−1) | B | 0.6×10-6 K−1 | −1 874 416.2 µm K | −1.1
Calculated length of the WP at 20 ∘C | l20 wp=125.0005 mm
Combined standard uncertainty | uc(l20 wp)=2.5 µm
Coverage factor k95=2 expanded uncertainty | U95=k95⋅uc(l20wp)=2⋅2.5 µm = 5.0 µm
(confidence level p=95 %)