## Worksheet 1
Error | Type | Standard | Sensitivity | Uncertainty
source |  | uncertainty | coefficient | contribution
xi |  | u(xi) | c(xi) | u(xi)⋅c(xi), µm
Calibrated length of the RWP at 20 ∘C
(l20 rwp=160.0013 mm) | B | 1.0 µm | 0.9 | 0.9
Measured length of the WP
(lwp=150.0015 mm) | B | 0.2 µm | 1 | 0.2
– random effects during measurements | A | 0.2 µm
– measuring instrument | B | 0.1 µm
Measured length of the RWP
(lrwp=160.0418 mm) | B | 0.2 µm | −0.9 | −0.2
– random effects during measurements | A | 0.2 µm
– measuring instrument | B | 0.1 µm
Measured temperature of the RWP
(trwp=Δtrwp+20 ∘C = 43.5 ∘C) | B | 0.6 K | 0 µm K−1 | 0
– random effects during measurements | A | 0.1 K
– measuring instrument | B | 0.6 K
Calibrated CTE of the RWP
(αrwp=23.6×10-6 K−1) | B | 0.6×10-6 K−1 | 0 µm K | 0
Known/assumed CTE of the scale
(αrwp=12.6×10-6 K−1) | – | Can be neglected | 0 µm K−1 | 0
Possible temperature difference between the RWP and the scale during measurements of length of the WP
(δtscwp=0 K) | B | 0.3 K | 1.9 µm K−1 | 0.5
Possible temperature difference between RWP and the scale temperature during measurements of length of the RWP
(δtscrwp=0 K) | B | 0.3 K | −1.9 µm K−1 | −0.5
Possible temperature difference between the RWP and the WP
(δtwp=0 K) | B | 0.3 K | −3.5 µm K−1 | −1.0
Possible CTE difference between RWP and WP
(δαwp=0×10-6 K−1) | B | 0×10-6 K−1 | −3 520 237.2 µm K | 0
Calculated length of the WP at 20 ∘C | l20 wp=149.9635 mm
Combined standard uncertainty | uc(l20 wp)=1.6 µm
Coverage factor k95=2 expanded uncertainty | U95=k95⋅uc(l20wp)=2⋅1.6 µm = 3.2 µm
(confidence level p=95 %)