## Worksheet 1
Quan-tity | Uncertainty source | Type of uncertainty estimation according to GUM | Relative standarduncertainty; σv,rel
fd; ftx | Random effects of Doppler frequency estimation, including phase jitter and electrical noise | Type A estimation from calibration measurements (Fig. 4) for (CRF 2D RF) and an SNR of 5 dB; normal distribution with σv,rel=0.4% | 0.4 %
fd; ftx | Unknown systematic effects of Doppler frequency estimation, including frequency-dependent attenuation of the fluid and drift in the slow-time clock source | Type B estimation from calibration measurements (Fig. 4) for an SNR of 5dB; uniformly distributed in the interval Δvrel=±0.08% | 0.05 %
c | Value of the speed of sound of the fluid cGaInSn=2740ms-1 (given by , without a measurement uncertainty) | Type A estimation based on ; σc,rel=0.03% | 0.03 %
| Unknown systematic variations of the speed of sound in the fluid due to temperature changes | Type B estimation for the sound–speed temperature coefficient of liquid gallium ; dcdϑ=-0.3m(sK)-1 for Δϑ=±10K rectangular distributed in the interval Δcrel=±0.11% | 0.06 %
v | Influence of the spatial resolution from the finite width of the sound field | Type B estimation for a beam width Δx=3 mm and typical velocity gradients of MHD experiments estimated from numerical simulation; dvdx=0.16mm-1⋅vmax ; Δv=1/2⋅dvdx⋅b Δvrel=±24% | 13.9 %
|  | Total uncertainty ∑iσv,rel,i2= | 13.9 %