## Worksheet 1
Uncertainty component related to |  | Fundamental | (2 to 150) kHz
|  | current | current harmonics
|  | Relative standard | Relative standard
|  | uncertainty | uncertainty
Acquisition card, | Calibration | 2.5×10-4 | 2.5×10-4
AI2, uUCS1/UCS1, | INL (integral nonlinearity) | 3.0×10-5 | 3.0×10-5
AI3, uUn,CS2,HPF/Un,CS2,HPF | DNL (diferential nonlinearity) | 1.5×10-5 | 1.5×10-5
Current sensor 1, | Calibration | 2.5×10-3 | 2.5×10-3 (up to 100 kHz)
uCCS1,AI2/CCS1,AI2 |  |  | 6.5×10-3 (from 100 to 150 kHz)
| Linearity with current | 6.3×10-4 | 8.9×10-3
| Linearity with frequency |  | 1.9×10-3
Current sensor 2 + HPF | Position influence | 4.8×10-3 | 4.8×10-3
uCCS2,HPF,AI3/CCS2,HPF,AI3 | Repeatability | 2.8×10-5 | 2.8×10-4
Daily drift |  | 8.2×10-6 | 1.5×10-4
Data treatment software | FFT windowing | 3.5×10-4 | 1.4×10-3
Noise | 1.5×10-4 | 3.1×10-4
Cable impedance |  | 4.4×10-7 | 1.7×10-4
Standard uncertainty (k=1) |  | 5.4×10-3 | 5.4×10-3
Expanded uncertainty (k=2) |  | 1.1×10-2 | 2.1×10-2 (up to 100 kHz)
3.2×10-2 (from 100 to 150 kHz)