## Worksheet 1
Radiator type |  |  |  | Plate radiator | Water-heat-pipe cavity radiator
Temperature radiator |  | tradiator | ∘C | 100 | 100
|  | Tradiator | K | 373.2 | 373.2
Radiance temperature |  | trad | ∘C | 100.2 | 100
|  | Trad | K | 373.4 | 373.2
uall(ΔTε) with rectangular distribution (W=1,7) | Uncertainty emissivity | δΔTε | ∘C | 0.23 | 0
u1 | ∘C | 0.13 | 0
Influence uncertainty ambient temperature | δTamb | ∘C | 0.01 | 0
|  | u2 | ∘C | 0 | 0
| Uncertainty spectral range (radiation thermometer) | δTλ, RT | ∘C | 0.06 | 0
|  | u3 | ∘C | 0.02 | 0
| Uncertainty spectral range (camera) | δTλ, C | ∘C | 0.02 | 0
|  | u4 | ∘C | 0.01 | 0
Calibration uncertainty radiation thermometer with normal distribution (W=1) |  | δΔT1 | ∘C | 0.04 | 0.04
|  | u5 | ∘C | 0.04 | 0.04
Uncertainty camera (Type A, noise) with normal distribution (W=1) |  | δTTyp A | ∘C | 0.05 | 0.05
|  | u6 | ∘C | 0.05 | 0.05
Temperature inhomogeneity (radiator) with rectangular distribution (W=1,7) |  | δTH | ∘C | 1 | 0.04
|  | u7 | ∘C | 0.58 | 0.02
Temporal stability of the temperature (radiator) with rectangular distribution (W=1,7) |  | δTZ | ∘C | 0.2 | 0.01
|  | u8 | ∘C | 0.12 | 0.01
Influence uncertainty SSE with rectangular distribution (W=1,7) |  | δTSSE | ∘C | 0.12 | 0.12
|  | u9 | ∘C | 0.07 | 0.07
Determined uncertainty |  | u (k=1) | ∘C | 0.61 | 0.1
|  | U (k=2) | ∘C | 1.22 | 0.19